## Slide 1: Connection Layout
Tiffany Liu 06/26/2023

## Slide 2
18 12 B1 A1 7 21 GND
17 B5* GND 10 22 6*
IDC- IDC+
IDC- IDC+
TF
STI
STM
STO
TOP
DOOR
TF
TF
STI
STM
STO
TOP
DOOR
TF
11
B3
1
B7
8
B2
15
B6 (large capacitance to B7 and B3)
14
16
A3
24
B4
3
A2
13
1
SR830
16
SR830
14
15 (open,ay)/23 (close,ax)
1
3
SR830
20
SR830
18
13(open,ay)/21
(close,ax)
3
Emitter
DAC
Breakout Box
DAC
Breakout Box
Collector
12 8
2, 23
STIBiasC = 2 (STI bias tee box)
STOBiasC = 22 ( STO bias tee box)
Filament top plate = 5, note: don’t use DAC port 17 (short to ground)
19,20
9
5
4
*when connecting do 2 with 20
SMALL GLASS DEWAR
STIBiasE = 11 (STI bias tee box)
STOBiasE= 24 ( STO bias tee box)
doorE/C_close = ax
doorE/C_open = ay
Aux 3,4=Vee, 10=Vee, Aux 1,2=Vdd
* Short to ground

## Slide 3
Emitter Side | Breakout Box | DAC
Top Metal | 8 | 14
STO | B7 | 24
STI | B3 | 11
STM | 1 | 16
Door | B2 | 15 (open)/23 (close)
Collector Side | Breakout Box | DAC
Top Metal | 3 | 18
STO | B4 | 22
STI | A3 | 2
STM | 24 | 20
Door | A2 | 13 (open)/21 (close)

## Slide 4
Grounding
  Nice to have a table top rack so instruments (i.e. metal surfaces) aren’t touching each other
  Earth ground
  USB not great, find usb to ethernet connector (cat)
Insulator
  charging
  Defects, tunneling,
  How this ties into large voltages
Think about using much smaller voltages, see how it looks like in walk on spheres calculation

## Slide 5
B1 2 B2 4 B3 6 18
B7 8 20 10 11 12
IDC- IDC+
IDC- IDC+
TF
STI
STM
STO
TOP
DOOR
TF
TF
STI
STM
STO
TOP
DOOR
TF
B4
7
B5
21
B6
9
23
19
5
A3
3
A2
1
A1
Emitter
DAC
Breakout Box
DAC
Breakout Box
Collector
12 8
14, 16
STIBiasC = 2 (STI bias tee box)
STOBiasC = 22 ( STO bias tee box)
Filament top plate DAC = 5
15,17
22
24
13
*when connecting do 14 with 15
BIG GLASS DEWAR
Aux 3,4=Vee, 10=Vee, Aux 1,2=Vdd
STIBiasE = 11 (STI bias tee box)
STOBiasE= 24 ( STO bias tee box)
doorE/C_close = ax
doorE/C_open = ay
3
SR830
20
SR830
18
13(open,ay)/21
(close,ax)
3
1
SR830
16
SR830
14
15(open,ay)/23(close,ax)
1

## Slide 6
B2 2 A2 4 B5 6 18
B7 8 20 10 11 12
IDC+ IDC-
IDC+ IDC-
TF
DOOR
TOP
STO
STM
STI
TF
B3
23
9
B6
21
A3
7
NA
Emitter
DAC
Breakout Box
DAC
Breakout Box
Collector
12 8
14, 15
16,17
24
13
*when connecting do 14 with 16
BIG GLASS DEWAR (RF)
1
15(open,ay)/23(close,ax)
14
SR830
16
SR830
1
TF
DOOR
TOP
STO
STM
STI
TF
B1
19
5
B4
3
A1
1
NA
3
13(open,ay)/21
(close,ax)
18
SR830
20
SR830
3
STIBiasC = 2 (STI bias tee box)
STOBiasC = 22 ( STO bias tee box)
22 (DAC=5)
Aux 3,4=Vee (-5V), 10=E (Gnd), Aux 1,2=Vdd (+5V)
doorE/C_close = ax
doorE/C_open = ay
STIBiasE = 11 (STI bias tee box)
STOBiasE= 24 ( STO bias tee box)

## Slide 7
B4 2 A2 4 B5 6 18
B7 8 10 11 12 A3
TFC IDC+ IDC- TFE
BiasD1 TFC IDC+ IDC- TFE BiasD2
DOOR OUT
TWIDDLE
TOP
STO
STM
STI
DOOR IN (MOD)
B3
23
9
B6
21
20
7
NA
Emitter
DAC
Breakout Box
DAC
Breakout Box
Collector
3 12 D1 8 D2 1
14, 15
16,17
24
13
*when connecting do 14 with 16
BIG GLASS DEWAR (RF)
15(open,ay)/23(close,ax)
17
14
SR830
19
SR830
9
DOOR OUT
TWIDDLE
TOP
STO
STM
STI
DOOR IN(MOD)
B1
19
5
B2
3
A1
1
NA
13(open,ay)/21
(close,ax)
7
18
SR830
20
SR830
10
STI BiasC = 2
STO BiasC = 22
22 (DAC=5)
Aux 3,4=Vee (-5V), E (Gnd), Aux 1,2=Vdd (+5V)
Aux1=+4.5V,Aux2=-1.5V, Aux3=+1.5V
doorE/C_close = ax
doorE/C_open = ay
STI BiasE= 4
STO BiasE= 24
11 3 S1 S2 1 16
*looking down at device

## Slide 8
B4 2 A2 4 B5 6 18
B7 8 10 11 12 A3
B3
23
9
B6
21
20
7
NA
Emitter
Breakout Box
Breakout Box
Collector
14, 15
16,17
24
13
BIG GLASS DEWAR (RF)
B1
19
5
B2
3
A1
1
NA
22 (DAC=5)
*looking down at device

## Slide 9
Emitter Side | Breakout Box | DAC
Top Metal | 9 | 14
STO | B6 | 24
STI | 20 | 4
STM | 21 | 19
Door Out | B3 | 15 (open)/23 (close)
TF | 12,18 | 1
Door In | 7 | 9
Collector Side | Breakout Box | DAC
Top Metal | 5 | 18
STO | B2 | 22
STI | A1 | 2
STM | 3 | 20
Door Out | B1 | 13 (open)/21 (close)
TF | 2,8 | 3
Door In | 1 | 10
 | BB | DAC
Sense | A3 | 16
Twiddle | 23 | 17
S2 | 11 | NA
D2 | 6 | NA
 | BB | DAC
Sense | B7 | 11
Twiddle | 19 | 7
S1 | 10 | NA
D1 | 4 | NA

## Slide 10: Notes on above
B4 has like a 40Mohm resistance to ground
  Idk how to fix that other than feeding new wire
TFs have ~10Mohm resistance to ground because they’re kinda capacitively coupled to nearby source and drain gates which have a resistance to ground
  Probably from residual flux in pin connectors??
  Strange, definitely not on wire side, but when disconnected from board, get quick OL between a TF and source/drain gate (neither shorted to ground when disconnected)
    But when connected, see the resistance to ground…
All resistances and capacitances read correctly, but not when I check shorts to ground
  A lot of pins have quick OL…not sure if this is a factor of the weather or the soldering of the pin connectors…
  See if it makes things weird when cold/during measurement, otherwise have ignored it

## Slide 11
Door out
Twiddle
TOP
STO
STM
STI
Sense
Door in
Door out
Twiddle
TOP
STO
STM
STI
Sense
Door in
TF C
TF C
IDC +
IDC +
IDC -
IDC -
TF E
TF E

## Slide 12
Emitter Side | Breakout Box | DAC
Top Metal | 9 | 14
STO | B6 | 24
STI | A3 | 11
STM | 21 | 16
Door | 23 | 15 (open)/23 (close)
TF | B3,7 | 1
Collector Side | Breakout Box | DAC
Top Metal | 5 | 18
STO | B4 | 22
STI | A1 | 2
STM | 3 | 20
Door | 19 | 13 (open)/21 (close)
TF | B1,1 | 3

## Slide 13: Wigner

## Slide 14
B2 2 A2 4 B5 6 18
B7 8 20 10 11 12
IDC- IDC+
IDC- IDC+
TF
STI
STM
STO
TOP
DOOR
TF
TF
STI
STM
STO
TOP
DOOR
TF
B3
7
A3
21
B6
9
23
19
5
B4
3
A1
1
B1
Emitter
DAC
Breakout Box
DAC
Breakout Box
Collector
12(DAC) 8(DAC)
SIM900 (ng,nf) SIM900 (ng,nf)
14, 15
BiasC = 2 (bias tee box)
Filament top plate DAC = 5, note: don’t use DAC port 17 (short to ground)
16,17
22
24
13
*when connecting do 14 with 16
BIG GLASS DEWAR (RF)
Aux 3,4=Vee, 10=Vee, Aux 1,2=Vdd
BiasE = 11 (bias tee box)
doorE/C_close = ax
doorE/C_open = ay
3
SR830 (current in)
SR830
SR830
18
13(open,ay)/21
(close,ax)
3
1
SR830 (sine out)
SR830
SR830
14
15(open,ay)/23(close,ax)
1

## Slide 15
B4 2 A2 4 B5 6 18 NA
B7 8 10 11 12 A3
B3
23
9
B6
21
20
7
NA
14, 15
16,17
24
13
B1
19
5
B2
3
A1
1
NA
22
*looking down at device